APPENDIX 5 — ACS712-20A CURRENT SENSOR WITH ESP32

AIM

To interface the ACS712-20A Hall-effect current sensor with the ESP32 ADC using a 10kΩ/20kΩ voltage divider and display the measured current (in amperes) on the Serial Monitor.

COMPONENTS REQUIRED

| Sl. | Component | Specification | Qty | Purpose |
| --- | --- | --- | --- | --- |
| 1 | Microcontroller | ESP32 Dev Board | 1 | ADC measurement |
| 2 | Current Sensor | ACS712-20A Module (±20A) | 1 | Measure current (Hall effect) |
| 3 | Resistor | 10 kΩ, 1/4 W | 1 | Voltage divider upper |
| 4 | Resistor | 20 kΩ, 1/4 W | 1 | Voltage divider lower |
| 5 | Test Load | Resistor/bulb (known current) | 1 | Current source for testing |

CIRCUIT DIAGRAM

ACS712-20A ESP32

VCC → 5V (VIN)

GND → GND

SOUT → [10kΩ] → node → GPIO34 (ADC)

node → [20kΩ] → GND

Max SOUT = 4.5V → V_node = 4.5 × 20/30 = 3.0V (safe)

Current path (series with load):

Supply (+) → IP+ ─────────────── IP- → Load → Supply(−)

THEORY

| Parameter | Value / Formula |
| --- | --- |
| Sensitivity (ACS712-20A) | 100 mV/A |
| Output at 0A | 2.500 V (half of 5V supply) |
| Divider ratio k | 20/(10+20) = 0.6667 |
| ADC voltage → SOUT voltage | V_out = V_adc / 0.6667 |
| SOUT → Current | I (A) = (V_out − 2.500) / 0.100 |
| Averaging | 500 ADC samples averaged to reduce noise |

PROGRAM CODE (ESP32 — Arduino IDE)

#define CURRENT_PIN 34

#define SAMPLES 500

#define DIVIDER_K 0.6667 // 20k/(10k+20k)

#define SENSITIVITY 0.100 // ACS712-20A: 100 mV/A

#define V_ZERO 2.500 // output at 0A

void setup() {

Serial.begin(115200);

analogReadResolution(12);

analogSetAttenuation(ADC_11db);

Serial.println("ACS712-20A Current Sensor");

}

void loop() {

long sum = 0;

for (int i = 0; i < SAMPLES; i++) {

sum += analogRead(CURRENT_PIN);

delayMicroseconds(100);

}

float adc_avg = (float)sum / SAMPLES;

float V_adc = (adc_avg / 4095.0) * 3.3;

float V_out = V_adc / DIVIDER_K;

float current = (V_out - V_ZERO) / SENSITIVITY;

Serial.print("ADC avg: "); Serial.print(adc_avg, 1);

Serial.print(" V_out: "); Serial.print(V_out, 3);

Serial.print(" V Current: "); Serial.print(current, 3);

Serial.println(" A");

delay(1000);

}

OBSERVATION TABLE

| Sl. | Load | ADC Avg | V_out (V) | Measured I (A) | Actual I (A) |
| --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |

PROCEDURE

1. Build the 10kΩ/20kΩ voltage divider: 10kΩ from SOUT to node; 20kΩ from node to GND; node to GPIO34.

2. Connect ACS712 VCC to 5V (VIN) and GND to GND.

3. Connect the test load in series through IP+ and IP− of the ACS712.

4. Upload the code. Open Serial Monitor at 115200 baud.

5. Power the load circuit and note the displayed current.

6. Verify with a known load (e.g., a resistor where I = V/R is known).

RESULT

RESULT The ACS712-20A current sensor was successfully interfaced with the ESP32. The SOUT output was level-shifted from 5V to 3.3V using a 10kΩ/20kΩ voltage divider connected to GPIO34. The code averaged 500 ADC samples, converted to actual voltage, reversed the divider, and calculated current using the 100mV/A sensitivity. Measured current values matched the expected values for known test loads.